# REQUEST FOR INFORMATION (RFI)

HIGH-CAPACITY CRYOGENIC STORAGE SYSTEMS WITH DEMONSTRATED LN₂ SUPPLY INDEPENDENCE AND EXTENDED OPERATIONAL AUTONOMY

Date of Request: 27 October 2026
RFI Reference: SAMRC RFI 26/04

The South African Medical Research Council (SAMRC) invites suitably qualified manufacturers, authorised distributors, service providers and other relevant market participants to provide information on commercially available high-capacity cryogenic storage systems that operate without connection to an external liquid nitrogen (LN₂) supply and without routine LN₂ replenishment.

Service providers are requested to provide the information specified in this RFI, including completion of the Functional Requirements Matrix in Appendix A, and the Commercial Response Schedule in Appendix B, together with relevant brochures, technical datasheets, application notes, validation/qualification documentation, reference information and other supporting evidence, where available.

Submission Contact:
Supply Chain Management South African Medical Research Council (SAMRC) Email: tenders@mrc.ac.za  Website: www.samrc.ac.za

Table of Contents

HIGH-CAPACITY CRYOGENIC STORAGE SYSTEMS WITH DEMONSTRATED LN₂ SUPPLY INDEPENDENCE AND EXTENDED OPERATIONAL AUTONOMY	4

1.	INTRODUCTION	4

2.	BACKGROUND	4

3.	PURPOSE	5

4.	DISCLAIMER	6

5.	REQUIRED INFORMATION	6

6.	HIGH-CAPACITY CRYOGENIC STORAGE SYSTEMS WITH DEMONSTRATED LN₂ SUPPLY INDEPENDENCE AND EXTENDED OPERATIONAL AUTONOMY	8

7.	ANALYTICAL AND APPLICATION CAPABILITY	8

8.	INFRASTRUCTURE, MONITORING, SAMPLE INTEGRITY AND INTEGRATION	9

8.1.	Infrastructure	9

8.2.	Monitoring and sample integrity	10

8.3.	Integration	10

9.	IMPLEMENTATION, COMMISSIONING AND TRAINING	10

10.	SUPPORT, MAINTENANCE AND LIFECYCLE	11

11.	PRICING AND COMMERCIAL INFORMATION	12

12.	MARKET INTELLIGENCE	12

13.	MARKET ASSESSMENT	13

14.	SUBMISSION REQUIREMENTS	13

15.	QUERIES AND SUBMISSION	14

16.	APPENDIX A – FUNCTIONAL REQUIREMENTS MATRIX	15

16.1.	Instructions	15

Platform information	15

16.2.	Functional requirements	15

16.3.	Declared performance and operational autonomy	17

17. APPENDIX B – COMMERCIAL RESPONSE SCHEDULE	18

17.1.	Instructions	18

17.2.  Commercial response	18

17.3. Commercial assumptions	19

17.4.  Total cost of ownership and operational cost disclosure	19

REQUEST FOR INFORMATION (RFI)

# HIGH-CAPACITY CRYOGENIC STORAGE SYSTEMS WITH DEMONSTRATED LN₂ SUPPLY INDEPENDENCE AND EXTENDED OPERATIONAL AUTONOMY

# INTRODUCTION

The South African Medical Research Council (SAMRC) invites suitably qualified manufacturers, authorised distributors, service providers, system integrators and other relevant market participants to provide information on commercially available high-capacity cryogenic storage systems capable of maintaining cryogenic storage conditions during normal operation without connection to an external liquid nitrogen (LN₂) supply and without routine LN₂ replenishment.

The RFI seeks solutions that eliminate dependence on routine LN₂ deliveries and continuous connection to a central LN₂ supply system, while maintaining the integrity of biological samples during supply interruptions, delivery delays, infrastructure failures or restricted site access.

Emphasis is placed on systems capable of maintaining all stored samples below −150°C for during normal operation, without connection to an external LN₂ supply and without routine LN₂ replenishment, supported by documented technical evidence and clearly defined operating conditions. Service providers shall identify whether their proposed systems require any LN₂ during normal operation, and if so in what volume and at what interval,, the technologies used to maintain storage temperatures, and any residual dependence on external supplies, infrastructure or operator intervention.

# BACKGROUND

The SAMRC supports health research and related scientific activities that require reliable long-term preservation of biological samples. Suitable cryogenic storage infrastructure is essential for maintaining sample integrity and supporting research continuity.

SAMRC is investigating the availability and suitability of high-capacity cryogenic storage systems that can provide increased storage capacity with a defined period of operational autonomy without external LN₂ replenishment.

For this RFI, the target application is dry vapour-phase cryogenic storage of biological samples, with a proposed minimum capacity of 30,000 cryovials, in which all stored samples remain below −150°C during normal operation without connection to an external LN₂ supply and without routine LN₂ replenishment.

In this RFI, dry vapour-phase storage means that samples are kept in nitrogen vapour above any liquid and are not in contact with liquid nitrogen. An external LN₂ supply means any source of liquid nitrogen outside the storage unit itself, including a central or bulk supply system, transfer piping, a supply dewar connected to the unit, a portable dewar used for filling, and an on-site nitrogen generator. LN₂ added for the initial charge, or to restore operation after a power failure or servicing, is not routine replenishment. Service providers shall declare the volume and frequency of any such additions.

Service providers should distinguish between systems that reduce routine LN₂ consumption, systems that can operate for defined periods without replenishment, and systems that can maintain storage conditions during interruptions to a central LN₂ supply. These capabilities shall be supported by clearly stated operating assumptions and technical evidence. Only systems that require no external LN₂ supply and no routine LN₂ replenishment during normal operation meet the mandatory requirement in Appendix A.

# PURPOSE

The purpose of this RFI is to obtain market information on commercially available high-capacity cryogenic storage systems and associated services relevant to SAMRC's requirements.

The information will be used to:

- Assess commercially available cryogenic storage technologies, configurations and market maturity.

- Determine the extent to which available systems meet the required storage capacity, temperature performance and operation without an external LN₂ supply.

- Understand LN₂ consumption, replenishment requirements, energy use and operating costs.

- Identify installation, infrastructure, monitoring, safety and data-management requirements.

- Assess implementation, commissioning, training, maintenance and lifecycle support arrangements.

- Obtain indicative acquisition costs and three-year and five-year total cost of ownership (TCO).

- Identify technically credible alternatives, limitations, supplier dependencies and operational risks.

# DISCLAIMER

This RFI is for information gathering purposes (market and price analysis) only and does not constitute a commitment to purchase, contract, or award work. Responses will not be evaluated as bids or proposals and will not form the basis of a contract or confer an advantage in any future procurement process. SAMRC may use the information gathered to refine its requirements, assess market capacity and competition, and determine the most appropriate procurement method. Participation is voluntary, and all response costs shall be borne by the service provider.

# REQUIRED INFORMATION

Service providers shall provide sufficient information to enable SAMRC to understand the capability, maturity, application suitability, implementation requirements, support model, lifecycle considerations, commercial model and market position of the proposed solution.

Complete Part A: Service provider Information and Part B: Service provider confirmation.

Detailed functional requirements shall be addressed in Appendix A, and detailed commercial information shall be provided in Appendix B.

  - Part A: Service provider Information

Legal name of organisation: __________________________________________

Manufacturer/platform proposed: ______________________________________

Service provider capacity: ☐ Manufacturer ☐ Authorised distributor ☐ Service provider ☐ System integrator ☐ Other

Country of registration: ______________________________________________

Company registration number, if applicable: _____________________________

South African presence: ☐ Yes ☐ No

Primary contact person: _____________________________________________

Position: _________________________________________________________

Telephone: _____________________________________________

Email: _________________________________________________

Website: _______________________________________________

Where applicable, service providers should provide evidence of manufacturer authorisation and describe their arrangements for local supply, technical support and maintenance.

  - Part B: Service provider confirmation

The service provider confirms that:

- The information submitted in response to this RFI is accurate to the best of the service provider’s knowledge.

- The person signing this schedule is authorised to submit the information on behalf of the service provider.

- The service provider understands that this RFI is issued for market information and planning purposes only.

- The response does not constitute a bid or proposal and will not result in the award of a contract.

- Participation in this RFI does not confer an advantage in any future procurement process.

- SAMRC may contact the service provider to request clarification or further information.

- All costs associated with preparing and submitting the response are borne by the service provider.

- Any information considered confidential has been clearly identified and the reasons for confidentiality have been stated.

Name of authorised representative: _______________________________________

Position or designation: ________________________________________________

Signature: __________________________________________________________

Date: ______________________________________________________________

# HIGH-CAPACITY CRYOGENIC STORAGE SYSTEMS WITH DEMONSTRATED LN₂ SUPPLY INDEPENDENCE AND EXTENDED OPERATIONAL AUTONOMY

For each proposed system or materially different configuration, service providers shall provide:

- Manufacturer, model and current commercially available version.

- Commercial availability and lifecycle status.

- System architecture and available configurations.

- Storage capacity, compatible cryovial and cryobox formats, and rack configuration.

- Storage technology, including whether the system uses dry vapour-phase storage.

- Whether any LN₂ is required during normal operation, and if so its volume and frequency.

- Maximum demonstrated autonomy without external LN₂ replenishment.

- Temperature performance during normal operation, sample access, LN₂ supply interruptions and loss of mains power.

- LN₂ consumption, replenishment intervals and operating assumptions.

- Electrical power requirements, backup power and UPS requirements.

- Monitoring, alarm and data-export capabilities.

- Installation requirements, accessories and associated services.

- Known limitations, dependencies and end-of-life considerations.

Detailed technical capability shall be provided in Appendix A.

# ANALYTICAL AND APPLICATION CAPABILITY

Service providers shall complete Appendix A and identify the extent to which the proposed solution supports relevant applications.

Service providers shall describe how the proposed system supports long-term storage of human- or animal-derived biological samples used in research, pre-clinical studies and clinical research.

Attention should be given to:

- Preservation of biological sample integrity at cryogenic temperatures.

- Storage of standard 2 mL cryovials (typical fill volumes of 1.0–1.5 mL).

- Accommodation of compatible 10 × 10 and 5 × 5 cryobox formats.

- Storage capacity and inventory organisation.

- Temperature uniformity and recovery following routine lid-opening events.

- Operational resilience during interruptions to LN₂ delivery, central supply, site access or mains power.

- Continuous monitoring, alarm notification and access to historical operating data.

- Evidence supporting the declared operation without routine LN₂ replenishment and temperature performance.

Service providers should distinguish between demonstrated performance, manufacturer specifications and projected performance. Any conditions or limitations affecting sample protection must be disclosed.

# INFRASTRUCTURE, MONITORING, SAMPLE INTEGRITY AND INTEGRATION

  - Infrastructure

Service providers shall identify:

- Electrical supply requirements compatible with South African 220–240 V power.

- Peak and steady-state power consumption.

- Heat rejection, ventilation and extraction requirements.

- Floor loading, dimensions, access and installation constraints.

- LN₂ supply, backup storage or other infrastructure requirements.

- UPS capacity expected runtime and supported functions during power interruptions.

- Requirements for safe operation in a mechanically ventilated laboratory.

The intended installation environment includes an existing cryogenic facility with a central LN₂ supply system, external bulk storage and vacuum-insulated transfer piping. Service providers shall state whether the proposed system requires connection to this central supply, or to any other external LN₂ supply, during normal operation.

  - Monitoring and sample integrity

Service providers shall describe:

- Continuous temperature and LN₂ level monitoring, where applicable.

- Alarm limits, local audible and visual alarms, and remote notification options.

- Monitoring-system access controls and configurable user roles, where supported.

- Export of time-stamped temperature logs, LN₂ levels, alarm events and relevant system status records.

- Available data formats and any proprietary software or license dependencies.

- Procedures for data backup, recovery and configuration restoration.

  - Integration

Service providers shall identify standard, optional, third-party or customised integration capabilities, including:

- Institutional monitoring or alarm systems.

- Laboratory information management systems (LIMS), where available.

- Remote monitoring and notification services.

- Data export through commonly used formats such as CSV, PDF, XML or equivalent.

All additional licenses, subscriptions, proprietary interfaces and integration costs shall be disclosed.

# IMPLEMENTATION, COMMISSIONING AND TRAINING

Service providers shall describe the proposed implementation approach, including:

- Site assessment and readiness requirements.

- Delivery, installation and commissioning.

- Factory acceptance testing (FAT) and site acceptance testing (SAT).

- Installation qualification (IQ) and operational qualification (OQ) support, where applicable.

- Verification of temperature performance, storage capacity, alarms and monitoring.

- Training for technical staff and researchers.

- Documentation, handover and operational support.

- Implementation duration, resource requirements, dependencies and risks.

The indicative implementation period is six to twelve weeks from the agreed commencement point, subject to confirmation of availability, site readiness and contractual arrangements.

# SUPPORT, MAINTENANCE AND LIFECYCLE

Service providers shall describe:

- Warranty coverage and exclusions.

- Preventive and corrective maintenance arrangements.

- Local or regional technical support and spare parts availability.

- Remote and on-site response arrangements.

- Availability of consumables and replacement components.

- Software and firmware updates, where applicable.

- Technical manuals, maintenance schedules and service reporting.

- Expected service life and lifecycle costs.

- Dependencies on proprietary parts, mandatory service contracts or third-party suppliers.

- The indicative support targets are remote technical response within 24–48 hours, on-site attendance within five business days for non-critical faults and within 48 hours for critical downtime faults, where locally available.

For planning purposes, the target support requirements are a minimum 12-month comprehensive warranty and a five-year service and maintenance plan. Service providers should indicate whether they can meet these requirements and identify any limitations.

# PRICING AND COMMERCIAL INFORMATION

Service providers shall complete Appendix B – Commercial Response Schedule and provide indicative pricing for:

- Base cryogenic storage system and required accessories.

- Components necessary to achieve the specified storage capacity and operation without an external LN₂ supply.

- Installation, commissioning and acceptance testing.

- Qualification support and training.

- Monitoring software, licenses and subscriptions.

- Warranty, maintenance and service contracts.

- Consumables and expected LN₂ usage.

- UPS, backup systems or other additional infrastructure.

- Shipping, insurance, duties, taxes and other applicable charges.

Service providers shall distinguish mandatory and optional costs and provide indicative Year 1, three-year and five-year TCO.

The commercial response should identify any costs necessary to operate without an external LN₂ supply or routine LN₂ replenishment, and quantify any LN₂ consumption and replenishment requirements under normal operating conditions

# MARKET INTELLIGENCE

Service providers are invited to provide evidence-based information on:

- Commercially available high-capacity cryogenic storage systems and suppliers.

- Market maturity and adoption.

- Technical advantages, limitations and competing storage technologies.

- Operational resilience and autonomy from routine LN₂ replenishment.

- Local and regional support availability.

- Energy consumption, LN₂ consumption and lifecycle cost drivers.

- Infrastructure, safety and operational risks.

- Monitoring, data management and interoperability.

- Vendor dependency and availability of alternative suppliers.

- Emerging technologies and relevant product developments.

Service providers should clearly distinguish between information about their own products, independently documented market information and their own assessments or opinions. Claims about competing systems should be supported by verifiable evidence wherever possible.

# MARKET ASSESSMENT

Service providers are requested to provide a concise assessment addressing:

- Available systems capable of meeting the proposed storage capacity and the requirement to operate without an external LN₂ supply or routine LN₂ replenishment.

- The extent to which a single system can satisfy all mandatory requirements.

- Alternative configurations or technologies that may provide equivalent or better resilience. Configurations that depend on an external LN₂ supply, as defined in the Background, do not meet the mandatory requirement.

- Requirements that are unsupported or only partially supported by available systems.

- Differences in LN₂ consumption, energy use, temperature performance and operating costs.

- Supplier concentration, local support and lifecycle risks.

- Infrastructure, installation and maintenance considerations.

- Major acquisition, implementation and lifecycle cost drivers.

- Risks associated with dependence on a single supplier or technology.

- Recommended risk controls and considerations for future procurement.

The assessment should not assume or predetermine the suitability of any manufacturer, system or supplier.

# SUBMISSION REQUIREMENTS

A complete response should include:

- Part A – Service provider Information.

- Part B – Service provider Confirmation.

- Proposed system overview and technical description.

- Functional and operational capability information.

- Infrastructure, monitoring and integration information.

- Implementation, qualification and training approach.

- Support, maintenance and lifecycle information.

- Market assessment and relevant references.

- Appendix A – Functional Requirements Matrix.

- Appendix B – Commercial Response Schedule.

- Supporting technical and commercial documentation.

Service providers should provide separate Appendix A and Appendix B schedules for materially different systems or configurations.

# QUERIES AND SUBMISSION

Queries concerning this RFI may be submitted to:

Supply Chain Management South African Medical Research Council (SAMRC) Email: tenders@mrc.ac.za

The closing date and time for submission are: 17 November 2026 at 11:00 AM (SAST)

# APPENDIX A – FUNCTIONAL REQUIREMENTS MATRIX

  - Instructions

Service providers shall complete this matrix for each proposed cryogenic storage system or materially different configuration, indicating support, availability, limitations, dependencies and associated costs.

### Platform information

- Service provider: ______________________________

- Manufacturer: _____________________________

- Platform/model: ___________________________

- Configuration: _____________________________

- Date: ____________________________________

  - Functional requirements

| No. | Requirement | Priority | Supported | Availability | Comments / evidence |
| --- | --- | --- | --- | --- | --- |
| 1a | During normal operation with mains power available, maintain all stored samples below −150°C without connection to an external LN₂ supply and without routine LN₂ replenishment. Declare any LN₂ addition required during normal operation, with its volume and frequency. | Mandatory |  |  |  |
| 1b | Provide cryogenic storage for at least 30,000 2mL cryovials | Mandatory |  |  |  |
| 2 | Provide dry vapour-phase storage, validated temperature performance at the warmest storage location, and stability during routine access. | Mandatory |  |  |  |
| 3 | Declare LN₂ consumption, refill triggers, replenishment intervals and operating assumptions. | Mandatory |  |  |  |
| 4 | Provide continuous temperature and applicable LN₂ level monitoring, with local audible/visual alarms and exportable temperature logs, alarm events and system records. | Mandatory |  |  |  |
| 5 | Provide remote alarm notifications or institutional alarm interfaces. | Mandatory |  |  |  |
| 6 | Declare electrical requirements, UPS capacity and functions maintained during power interruptions. | Mandatory |  |  |  |
| 7 | Supply compatible stainless-steel or aluminium racks for 10 × 10 and 5 × 5 cryoboxes, with a rack map showing positions, formats and total vial capacity. | Mandatory |  |  |  |
| 8 | Operate at ambient temperatures of 10–40°C or clearly disclose any limitations. | Preferred |  |  |  |
| 9 | Provide installation, commissioning, performance verification, and applicable FAT, SAT and IQ/OQ support and documentation. | Mandatory |  |  |  |
| 10 | Provide a minimum 12-month comprehensive warranty, a five-year service and maintenance plan, local technical support, response times and spare parts availability. | Mandatory |  |  |  |
| 11 | Support normal offline operation without network connectivity. | Mandatory |  |  |  |
| 12 | Provide password-protected access and configurable user roles. | Mandatory |  |  |  |
| 13 | Support LIMS integration and remote monitoring. | Preferred |  |  |  |
| 14 | Declare LN₂ consumption and electrical power consumption under the stated operating conditions. | Mandatory |  |  |  |
| 15 | Provide oil-free cooling or demonstrate equivalent reliability, maintainability and sample protection. | Mandatory |  |  |  |
| 16 | Provide on-site training and materials covering operation, maintenance and troubleshooting. | Mandatory |  |  |  |
| 17 | Supply technical manuals, maintenance schedules, sensor verification documentation and applicable validation records. | Mandatory |  |  |  |

  - Declared performance and operational autonomy

Service providers shall complete the following table and identify the assumptions used to determine each declared value.

| Performance parameter | Unit | Service provider's declaration |  |
| --- | --- | --- | --- |
| Volume and frequency of any LN₂ addition during normal operation. |  | L; days |  |
| Holding time below −150°C after loss of mains power, with no LN₂ added. |  | Days |  |
| Duration below −150°C during LN₂ interruption. | Days |  |  |
| LN₂ consumption and replenishment interval. | L/day; days |  |  |
| Assumed ambient temperature and storage load. | °C; % capacity |  |  |
| Steady-state and peak electrical power | W |  |  |
| Recommended UPS capacity and runtime | VA/W; minutes |  |  |
| Functions maintained during power interruptions | Description |  |  |
| Connection to an external LN₂ supply required during normal operation | Yes / No |  |  |
| Temperature recovery after lid opening | Minutes |  |  |
| Installation and commissioning requirements | Description |  |  |
| Maintenance and service arrangements | Description |  |  |

Evidence for declared performance: Service providers shall explain how operation without routine LN₂ replenishment, and the holding time after loss of mains power, were established, including the initial operating condition, LN₂ charge or other stored cooling capacity, storage load, ambient temperature, lid-opening frequency and any required operator intervention. Manufacturer validation data or documented performance test results should be supplied where available.

## 17. APPENDIX B – COMMERCIAL RESPONSE SCHEDULE

### 17.1.	Instructions

Service providers shall provide indicative pricing for commercially available, high-capacity cryogenic storage systems that operate without an external liquid nitrogen (LN₂) supply and without routine LN₂ replenishment. Pricing shall distinguish mandatory components from optional features and services and support comparison of the estimated three-year and five-year total cost of ownership (TCO).

All prices are requested for market research and procurement planning only and do not constitute a quotation or commitment to purchase. Service providers shall identify all costs, assumptions, exclusions and dependencies, particularly those affecting operation without an external LN₂ supply.

### 17.2.  Commercial response

Service providers shall complete the following schedule in the nominated currency. Costs required to achieve operation without an external LN₂ supply, and the stated temperature performance must be clearly identified.

| Cost element | Requirement | Once-off cost | Annual cost | 3-year TCO | 5-year TCO |
| --- | --- | --- | --- | --- | --- |
| Base cryogenic storage system meeting the mandatory requirements in Appendix A | Mandatory |  |  |  |  |
| Racks, cryobox compatibility and accessories | Mandatory |  |  |  |  |
| Monitoring, temperature control and alarms | Mandatory |  |  |  |  |
| LN₂ management or backup systems, where required | Mandatory |  |  |  |  |
| Software, licenses and subscriptions | As applicable |  |  |  |  |
| Installation, commissioning and FAT/SAT | Mandatory |  |  |  |  |
| IQ/OQ qualification, where applicable | Mandatory |  |  |  |  |
| Training and application support | Mandatory |  |  |  |  |
| Warranty and service/maintenance contracts | Mandatory |  |  |  |  |
| Consumables and routine operating supplies | As applicable |  |  |  |  |
| Third-party products and services | As applicable |  |  |  |  |
| Shipping, insurance, duties and taxes | As applicable |  |  |  |  |
| Optional accessories and enhanced monitoring | Optional |  |  |  |  |
| Other costs (specify) | As applicable |  |  |  |  |
| Total estimated cost |  |  |  |  |  |

TCO assumptions: Year 1 costs shall be included in the three-year and five-year TCO calculations. Service providers shall identify whether each cost is once-off or recurring and avoid double counting costs included elsewhere.

### 17.3. Commercial assumptions

- Currency: ______________________________________________

- Price validity period: ____________________________________

- VAT: ☐ Included ☐ Excluded

- Shipping and insurance: ☐ Included ☐ Excluded ☐ Separate

- Import duties and customs: ☐ Included ☐ Excluded ☐ N/A

- Other taxes or charges: __________________________________

- Annual price escalation: _________________________________

- Mandatory annual maintenance and service: ________________

- Mandatory licenses and subscriptions: _____________________

- Estimated annual LN₂ consumption and cost assumptions:

- Other assumptions, exclusions or dependencies:

### 17.4.  Total cost of ownership and operational cost disclosure

Service providers shall briefly explain the assumptions used to calculate the three-year and five-year TCO, including:

- Initial acquisition, delivery, installation and commissioning costs.

- Annual maintenance, servicing, licenses, subscriptions and consumables.

- Expected LN₂ consumption and replenishment requirements under normal operating conditions.

- Any additional equipment, infrastructure, backup systems or services required to operate without an external LN₂ supply.

- Annual price escalation, warranty coverage, exclusions and anticipated replacement costs.

Important: All costs must be presented consistently, with once-off and recurring expenses clearly distinguished. Any optional costs or dependencies necessary to achieve the stated storage capacity, temperature or operation without an external LN₂ supply or routine LN₂ replenishment must be disclosed separately. The SAMRC will use this information for market analysis, affordability assessment and future procurement planning.